NIST Reference Standard run on 5-16-2021. Peak positions and relative intensities are within the expected range.

## Measurement Conditions:  (Bookmark 1)

Dataset Name	NIST_Reference_05_16_2021

File name	C:\XRD Data\NIST_Reference\NIST_Reference_05_16_2021.xrdml

Comment	Configuration=Sample Spinner, Owner=User-1, Creation date=11/1/2006 2:23:59 PM

Goniometer=PW3050/60 (Theta/2Theta); Minimum step size 2Theta:0.001; Minimum step size Omega:0.001

Sample stage=Spinner PW3064

Diffractometer system=XPERT-PRO

Measurement program=C:\PANalytical\Data Collector\Programs\NIST_Reference.xrdmp, Identifier={937D76BB-8184-4FFD-B032-2B4C339E7253}

PHD Lower Level = 6.28 (keV), PHD Upper Level = 12.80 (keV)

Measurement Start Date/Time	5/16/2021 9:05:28 AM

Operator	User-1

Raw Data Origin	XRD measurement (*.XRDML)

Scan Axis	Gonio

Start Position [°2θ]	20.0042

End Position [°2θ]	139.9962

Step Size [°2θ]	0.0080

Scan Step Time [s]	10.1600

Scan Type	Continuous

PSD Mode	Scanning

PSD Length [°2θ]	2.12

Offset [°2θ]	0.0000

Divergence Slit Type	Fixed

Divergence Slit Size [°]	0.2500

Specimen Length [mm]	10.00

Measurement Temperature [°C]	25.00

Anode Material	Cu

Intended Wavelength Type	K-Alpha

K-Alpha1 [Å]	1.54060

Generator Settings	40 mA, 45 kV

Diffractometer Type	0000000067011063

Diffractometer Number	0

Goniometer Radius [mm]	240.00

Dist. Focus-Diverg. Slit [mm]	100.00

Incident Beam Monochromator	No

Spinning	No

## Main Graphics, Analyze View: (Bookmark 2)

## Peak List: (Bookmark 3)

| Pos. [°2θ] | Height [cts] | FWHM Left [°2θ] | d-spacing [Å] | Rel. Int. [%] |
| --- | --- | --- | --- | --- |
| 25.6935 | 671.89 | 0.4608 | 3.46445 | 24.06 |
| 35.2648 | 2792.89 | 0.4608 | 2.54301 | 100.00 |
| 37.9216 | 214.04 | 0.3072 | 2.37073 | 7.66 |
| 41.8178 | 134.08 | 0.4608 | 2.15841 | 4.80 |
| 43.4778 | 987.33 | 0.4608 | 2.07977 | 35.35 |
| 52.6700 | 574.92 | 0.4608 | 1.73640 | 20.59 |
| 57.6177 | 2372.45 | 0.4608 | 1.59850 | 84.95 |
| 59.8795 | 16.25 | 0.4608 | 1.54341 | 0.58 |
| 61.4117 | 536.73 | 0.4608 | 1.50853 | 19.22 |
| 66.6452 | 355.27 | 0.4608 | 1.40219 | 12.72 |
| 68.3388 | 326.21 | 0.4608 | 1.37151 | 11.68 |
| 70.4982 | 5.59 | 0.9216 | 1.33470 | 0.20 |
| 74.4509 | 51.38 | 0.4608 | 1.27332 | 1.84 |
| 77.0640 | 935.23 | 0.4608 | 1.23653 | 33.49 |
| 80.7805 | 38.16 | 0.3072 | 1.18875 | 1.37 |
| 84.4814 | 44.95 | 0.4608 | 1.14586 | 1.61 |
| 86.5671 | 81.68 | 0.3072 | 1.12352 | 2.92 |
| 89.1072 | 379.90 | 0.4608 | 1.09795 | 13.60 |
| 90.8189 | 250.62 | 0.4608 | 1.08167 | 8.97 |
| 95.3548 | 218.24 | 0.4608 | 1.04184 | 7.81 |
| 98.5450 | 11.86 | 0.4608 | 1.01647 | 0.42 |
| 101.1753 | 441.31 | 0.4608 | 0.99703 | 15.80 |
| 103.4128 | 17.75 | 0.4608 | 0.98147 | 0.64 |
| 111.0903 | 54.64 | 0.4608 | 0.93418 | 1.96 |
| 114.1663 | 47.85 | 0.4608 | 0.91761 | 1.71 |
| 116.6882 | 534.36 | 0.4608 | 0.90494 | 19.13 |
| 117.9319 | 54.61 | 0.3072 | 0.89898 | 1.96 |
| 122.1032 | 38.55 | 0.4608 | 0.88029 | 1.38 |
| 124.6682 | 43.09 | 0.4608 | 0.86974 | 1.54 |
| 127.7637 | 385.83 | 0.4608 | 0.85790 | 13.81 |
| 129.9568 | 323.35 | 0.3072 | 0.85008 | 11.58 |
| 131.1725 | 675.07 | 0.3072 | 0.84594 | 24.17 |
| 136.1485 | 285.10 | 0.4608 | 0.83036 | 10.21 |

## Pattern List: (Bookmark 4)

| Visible | Ref.Code | Score | Compound Name | Displ.[°2θ] | Scale Fac. | Chem. Formula |
| --- | --- | --- | --- | --- | --- | --- |
| * | 04-013-1687 | 72 | Aluminum Oxide | 0.116 | 0.454 | Al2 O3 |