This is the simple example template containing only headers for each report item and the bookmarks. The invisible bookmarks are indicated by text between brackets.

Modify it according to your own needs and standards.

## Measurement Conditions:  (Bookmark 1)

Dataset Name	NIST_reference_05162021

File name	C:\XRD Data\NIST Reference\NIST_reference_05162021.xrdml

Comment	Configuration=Reflection-transmission spinner, Owner=User-1, Creation date=8/18/2015 10:52:11 AM

Goniometer=Theta/Theta; Minimum step size 2Theta:0.0001; Minimum step size Omega:0.0001

Sample stage=Reflection-transmission spinner; Minimum step size Phi:0.1

Diffractometer system=EMPYREAN

Measurement program=C:\PANalytical\Data Collector\Programs\NIST_reference.xrdmp, Identifier={674AD157-CDDC-4198-9C48-A1F8E2BF927F}

PHD Lower Level = 4.02 (keV), PHD Upper Level = 11.26 (keV)

Measurement Start Date/Time	5/16/2021 9:37:09 AM

Operator	tavakoli

Raw Data Origin	XRD measurement (*.XRDML)

Scan Axis	Gonio

Start Position [°2θ]	20.0066

End Position [°2θ]	139.9836

Step Size [°2θ]	0.0130

Scan Step Time [s]	8.6700

Scan Type	Continuous

PSD Mode	Scanning

PSD Length [°2θ]	3.35

Offset [°2θ]	0.0000

Divergence Slit Type	Fixed

Divergence Slit Size [°]	0.1089

Specimen Length [mm]	10.00

Measurement Temperature [°C]	25.00

Anode Material	Cu

K-Alpha1 [Å]	1.54060

Generator Settings	40 mA, 45 kV

Diffractometer Type	0000000011171041

Diffractometer Number	0

Goniometer Radius [mm]	240.00

Dist. Focus-Diverg. Slit [mm]	100.00

Incident Beam Monochromator	No

Spinning	No

## Main Graphics, Analyze View: (Bookmark 2)

## Peak List: (Bookmark 3)

| Pos. [°2θ] | Height [cts] | FWHM Left [°2θ] | d-spacing [Å] | Rel. Int. [%] |
| --- | --- | --- | --- | --- |
| 25.6687 | 150.15 | 0.6140 | 3.46774 | 25.20 |
| 35.2525 | 595.83 | 0.6140 | 2.54386 | 100.00 |
| 37.8789 | 21.24 | 0.8187 | 2.37330 | 3.56 |
| 41.7710 | 32.43 | 0.6140 | 2.16072 | 5.44 |
| 43.5762 | 86.74 | 1.2280 | 2.07530 | 14.56 |
| 52.6565 | 132.52 | 0.6140 | 1.73681 | 22.24 |
| 57.6039 | 214.95 | 1.2280 | 1.59885 | 36.08 |
| 59.8309 | 2.82 | 0.6140 | 1.54455 | 0.47 |
| 61.4018 | 121.10 | 0.6140 | 1.50874 | 20.32 |
| 66.6220 | 70.16 | 0.6140 | 1.40263 | 11.77 |
| 68.3033 | 68.75 | 0.6140 | 1.37214 | 11.54 |
| 74.3933 | 13.44 | 0.6140 | 1.27416 | 2.26 |
| 77.0676 | 250.58 | 0.6140 | 1.23648 | 42.06 |
| 80.7356 | 8.52 | 0.6140 | 1.18930 | 1.43 |
| 84.2518 | 3.87 | 0.8187 | 1.14839 | 0.65 |
| 86.5560 | 15.23 | 0.6140 | 1.12364 | 2.56 |
| 89.0843 | 58.04 | 0.6140 | 1.09818 | 9.74 |
| 90.9076 | 36.69 | 0.6140 | 1.08084 | 6.16 |
| 95.3352 | 35.92 | 0.6140 | 1.04200 | 6.03 |
| 98.3020 | 3.70 | 0.6140 | 1.01833 | 0.62 |
| 101.1645 | 66.19 | 0.6140 | 0.99710 | 11.11 |
| 103.1299 | 4.18 | 0.6140 | 0.98339 | 0.70 |
| 111.0822 | 5.55 | 0.6140 | 0.93423 | 0.93 |
| 116.7023 | 43.61 | 0.8187 | 0.90487 | 7.32 |
| 122.1837 | 5.62 | 1.2280 | 0.87994 | 0.94 |
| 124.6123 | 6.70 | 0.6140 | 0.86996 | 1.12 |
| 127.7320 | 27.89 | 0.6140 | 0.85802 | 4.68 |
| 129.8261 | 40.07 | 0.6140 | 0.85053 | 6.72 |
| 135.8337 | 54.34 | 0.6140 | 0.83128 | 9.12 |

## Pattern List: (Bookmark 4)

| Visible | Ref.Code | Score | Compound Name | Displ.[°2θ] | Scale Fac. | Chem. Formula |
| --- | --- | --- | --- | --- | --- | --- |
| * | 01-071-1683 | 40 | Aluminum Oxide | 0.038 | 0.199 | Al2 O3 |

## Document History: (Bookmark 5)

Beam attenuation factor(s) used to back-calculate ESDs:

- Modification time = "5/16/2021 10:02:44 AM"

- Modification editor = "tavakoli"

Insert Measurement:

- File name = "NIST_reference_05162021.xrdml"

- Modification time = "5/16/2021 10:02:44 AM"

- Modification editor = "tavakoli"

Default properties:

- Measurement step axis = "None"

- Internal wavelengths used from anode material: Copper (Cu)

- Original K-Alpha1 wavelength = "1.54060"

- Used K-Alpha1 wavelength = "1.54060"

- Original K-Alpha2 wavelength = "1.54443"

- Used K-Alpha2 wavelength = "1.54443"

- Original K-Beta wavelength = "1.39225"

- Used K-Beta wavelength = "1.39225"

- Irradiated length = "10.00000"

- Spinner used = "No"

- KBeta filter material = "Ni"

- KBeta filter thickness = "0.02000"

- Distance to sample = "Diffracted radius"

- Step axis value = "0.00000"

- Offset = "0.00000"

- Sample length = "10.00000"

- Modification time = "5/16/2021 10:02:44 AM"

- Modification editor = "tavakoli"

Interpolate Step Size:

- Initial Scan Range = 20.00660 - 139.99100

- Initial Step Size = 0.01313

- Derived Step Size = 0.01300

- Use Derived Step Size = "Yes"

- Modification time = "5/16/2021 10:02:44 AM"

- Modification editor = "PANalytical"

Determine Background:

- Add to net scan = "Nothing"

- User defined intensity = "0"

- Correction method = "Automatic"

- Bending factor = "2"

- Minimum significance = "1"

- Minimum tip width = "0.01"

- Maximum tip width = "1"

- Peak base width = "2"

- Use smoothed input data = "Yes"

- Granularity = "52"

- Modification time = "5/13/2021 4:58:24 PM"

- Modification editor = "tavakoli"

Search Peaks:

- Minimum significance = "0.75"

- Minimum tip width = "0.5"

- Maximum tip width = "5"

- Peak base width = "6"

- Method = "Minimum 2nd derivative"

- Modification time = "5/13/2021 5:01:07 PM"

- Modification editor = "tavakoli"

Search & Match:

- Allow pattern shift = "Yes"

- Auto residue = "Yes"

- Data source = "Profile and peak list"

- Demote unmatched strong = "Yes"

- Multi phase = "Yes"

- Restriction set = "Default"

- Restriction = "None"

- Subset name = ""

- Match intensity = "Yes"

- Two theta shift = "0"

- Identify = "No"

- Max. no. of accepted patterns = "5"

- Minimum score = "30"

- Min. new lines / total lines = "60"

- Search depth = "10"

- Minimum new lines = "5"

- Minimum scale factor = "0.1"

- Intensity threshold = "0"

- Use line clustering = "Yes"

- Line cluster range = "1.5"

- Search sensitivity = "1.8"

- Use adaptive smoothing = "Yes"

- Smoothing range = "1.5"

- Threshold factor = "3"

- Match Threshold = "0"

- N * Esds = "-1"

- Raw Weight = "-1"

- Peak Shape = "-1"

- Accepted Shape = "-1"

- Peak Power = "-1"

- New Peak Power = "-1"

- Intensity Power = "-1"

- N Peaks Power = "-1"

- Modification time = "5/12/2021 4:35:57 PM"

- Modification editor = "tavakoli"

Strip K-α2:

- Method = "Rachinger"

- Intensity ratio = "0.5"

- Delta wavelength ratio = "0"

- Modification time = "5/16/2021 10:04:52 AM"

- Modification editor = "PANalytical"

Delete Peak(s):

- Start position = "114.1642"

- End position = "114.1642"

- Modification time = "5/16/2021 10:05:11 AM"

- Modification editor = "tavakoli"

More items… (Bookmark 6)

More items… (Bookmark 7)

More items… (Bookmark 8)

More items… (Bookmark 9)

More items… (Bookmark 10)

More items… (Bookmark 11)

More items… (Bookmark 12)

More items… (Bookmark 13)

More items… (Bookmark 14)

More items… (Bookmark 15)